## Slide 1: A nonwoven whose reagent lines react to urine
UTIS TECHNICAL BRIEF | SAMPLE FOR THE IMH SMART PRESSURE-INJURY CARE R&D
In-house formulated water-soluble reagents are dispensed as lines on a moving nonwoven web, dried, and supplied as roll stock.
The sample (schematic)
Nonwoven roll
Dried reagent lines
Wetted zone: colour per analyte
Line colours are illustrative; the UTIS marker panel is not yet fixed.
How colour appears
General principle
1 Wicking: urine spreads along fibres and pores
2 Rehydration: the dried reagent partly dissolves or hydrates
3 Reaction: analyte yields colour or shifts an indicator [1]
4 Reading: by eye or camera, vs a UTIS colour chart (to be set)
Typical reagent-strip chemistries (illustrative)
General principle [1]
Nitrite
Diazotisation + coupling in acid: pink
Leukocyte esterase
Esterase hydrolysis + diazo coupling: purple
pH
Indicator acid/base form; colours vary
Protein
Buffered indicator binds protein: yellow to green
Principles of commercial urine strips. The UTIS composition, marker panel and read time are not yet fixed.
Early bench check
Developer-reported
With the same standard solution, the coated nonwoven matched the dipstick comparator in colour and reaction.
Tested conditions only; not an accuracy claim. Comparator, n, concentrations, read time and raw data to be confirmed.
[1] Siemens urine reagent-strip instructions for use · Source: UTIS TS 001 v0.1 (30 Sep 2026)
Notes: UTIS is a nonwoven web carrying dried reagent lines. In the general reagent-strip model, urine rehydrates the dried reagent and each line reacts with its analyte: nitrite and leukocyte-esterase lines colour only above their detection level, pH always takes a colour, and protein shifts from yellow towards green. UTIS's actual reactions still have to be confirmed with its own formulation, and a UTIS-specific colour chart is still to be set because the nonwoven background changes the colour seen. The four chemistries are the general principles of commercial urine strips, not a confirmed UTIS panel; the protein pad is mainly albumin-sensitive and is not an albumin-specific assay. Colour intensity should not be assumed proportional across all ranges, nor stable over time. The bench comparison was reported by the development team; test counts, concentrations, comparator type and raw images will be shared once provided and confirmed.

## Slide 2: Made roll-to-roll; scale-up under review
MANUFACTURING
The web is unwound, coated with reagent lines, dried and rewound as raw-material rolls for conversion into finished products.
Process (steps 1–5): developer-reported
Step 6 and control items: recommended
1 Unwind
Web spec, tension, wrinkles, speed
2 Reagent prep
Batch record, homogeneity, pot life
3 Nozzle coating
Dose, line width, position, spacing
4 Drying
Web temperature, dwell time, residual moisture
5 Rewind
Residual heat, roll pressure, set-off
6 Inspect & release
Reaction and uniformity at start, middle, end
Formulation must
Keep reactivity
Hold line shape
Anchor when dry
Stay put when wet (in use)
Stay stable in storage
Components, ratios, nozzle and drying settings are held in an internal manufacturing standard (confidential); step control items will be finalised there.
Dipstick vs UTIS
UTIS column: developer-reported
 | Typical dipstick [2] | UTIS
Carrier | Absorbent test pad | Nonwoven roll
Reagent | Impregnated, then dried | In-house mix, nozzle-dispensed
After drying | Pad fixed on a stick | Rewound as roll stock
Design focus | Discrete pad, set read time | Spread, reactivity, run speed
Output, metres of reagent web per day (approx.)
Chart: 
| Today | Target (assumed)
m/day | 3000 | 100000
×33
only if target verified
Today: developer-reported
Target: unverified estimate
Unverified: drying, line width and reactivity at speed; per-day basis to be confirmed. Product yield depends on web width, lanes and cut length.
[2] US FDA 510(k) K160372 decision summary (typical dipstick pad construction only)
Notes: The line unwinds the nonwoven, dispenses the in-house reagent mix through nozzles, dries the web on hot-air and hot-plate sections and rewinds it as roll stock; this process description comes from the development team. The control items under each step are recommended and will be finalised in the internal manufacturing standard. Formulation details stay confidential; what we share is what each variable controls. Current output of about 3,000 m per day is developer-reported. About 100,000 m per day after line modification is an unverified estimate: drying, line-width variation and reactivity at higher speed still have to be checked, and the per-day basis of the target is to be confirmed.

## Slide 3: Where it could fit, and what we need to test
PRESSURE-INJURY CARE FIT AND VALIDATION
Moisture and incontinence are recognised pressure-injury risk factors [5]. Two concepts, open questions and an analytical test path.
Pad / underpad insert
Reagent web in a pad or underpad, reagent side away from skin; lines read urine chemistry, not wetness.
Concept: current web, untested format
Dressing-integrated indicator
Indicator zone in a dressing, off the wound bed (e.g. exudate pH). Needs its own chemistry and safety tests.
Exploratory: not yet tested
Open technical questions
Line cross-talk
Wet reagent may migrate and alter the next line's colour or chemistry. Test spacing under single and repeat voids; no barriers yet.
Read window
Developer-reported: colour may shift after ~4 h (start point not stated) vs 4–5 h product changes on site. Stability is unproven.
Repeatability and skin safety
Lot, day, operator and web-position effects, and reagent leaching onto skin when wet, are not yet assessed.
Proposed analytical validation path
Proposal, not a fixed spec
1m
5m
15m
30m
1h
2h
4h
5h
6h
Record colour and result from first urine contact. Amber: reported shift after ~4 h (time zero unconfirmed).
Agreement across concentrations
Lot, day, operator repeatability
Repeat voids, cross-talk to 6 h
Real urine and finished product
Storage, interference (vitamin C)
Binary: CLSI EP12 [3]
pH, graded: own plan [4]
Bench agreement is not a test certificate, regulatory approval or clinical (e.g. UTI) diagnostic performance.
Proposed next steps with IMH
Agree target markers and output format
Co-design test protocol and acceptance criteria
MONIT supplies sample sets (lots, layouts)
Use images from the validated read window for AI reading (timing to agree with IMH)
[3] CLSI EP12, 3rd ed. · [4] CLSI method verification · [5] EPUAP/NPIAP/PPPIA guideline, 2019
Notes: Concept A places the current reagent web in a flat pad or underpad, reagent side away from the skin. Its lines report urine chemistry, not wetness, and this format has not been built or tested. Moisture and incontinence are recognised pressure-injury risk factors (EPUAP/NPIAP/PPPIA International Guideline, 2019); how the indicator could support that care is a question to scope with IMH. Concept B, a dressing-integrated indicator such as exudate pH, is exploratory and needs its own chemistry and safety testing. Open points: cross-talk between adjacent lines and the effect of single versus repeated voids; the read window (the development team reports colour may shift after about four hours, without stating the start point, while products at the site they observed are changed every four to five hours; time zero should be first urine contact, and stability both before four hours and up to the change is not yet shown); repeatability; and reagent leaching onto skin. The time points and evaluation items are a development proposal, not a fixed specification. This is bench analytical testing only; no clinical or UTI diagnostic performance, certification or regulatory status is claimed.